## Slide 1: Fraction collector behavior – ÄKTA pure F9-C and ÄKTA avant built in fraction collector
Pre-requisite | What to do | Result
A manual or method run is on-going | Execute the instruction ’Reset tube type’ and select which type you want to reset NOTE! Only enabled when no fractionation is on-going | Frac position will be reset for selected tube type(s) and collection will start in FIRST position next time fractionation is started
System is in Pause caused by out of tubes | Open hatch and replace filled tubes Close hatch (fraction collector performs a scan) then continue | Frac position will be reset for the current tube type and collection will continue in FIRSTposition
How can I get it to start collecting in first position?

## Slide 2
Fraction collector behavior – ÄKTA pure F9-C and ÄKTA avant built in fraction collector
Where will it start collecting when a fractionation is issued?
Pre- requisite | Action | Start in
System in End state | Start a new run and start fractionation | Next position
System running, no fractionation on-going | Open hatch (system will pause) Close hatch and continue | Next position
System running and fractionation is on-going | Open hatch (system will pause) Close hatch and continue | Current position
System is in Pause (not caused by out of tubes), fractionation is on-going | Open hatch Close hatch and continue | Current position
System is in Pause caused by out of tubes | Open and close hatch, then continue | First position